# Load Fridge with Cryostat Cold

## Fridge Preparation

### Check oil in the pumps (needs to be done once a month). If levels are low, need to refill.

### With fridge in mobile support frame, pump dilution fridge lines (still and condenser) overnight.

##### Connect T-piece to both condenser and still lines at the same time.

##### Connect an external turbo pump to the T-piece

##### Start in roughing pump mode

###### First use arrows < > to navigate to mode 023:Motor TMP off (press both arrows at the same time to confirm your selection)

##### Pump the pumping line first (with all fridge valves closed).

###### Pump for ~5mins until reading on turbo pump is <1x10-3 mbar.

###### Open the gate valve to the Still.

###### Open the S/V to the Condenser

(Pressure should rise to ~9x10-1mbar then drop back down).

##### After a few minutes, switch to turbo pump mode.

##### Leave to run overnight.

###### (Pressure in turbo pump gauge should drop from ~10-2 mbar to ~1.4x1-4 mbar)

##### Close all fridge valves and stop the pump.

##### Wait until pumping line is vented before disconnecting the pump and T-piece, put plastic caps on fridge line flanges.

### Clean both sides of IVC indium seal with a light dish scourer, working so that scratches on the seal are concentric rather than radial. Wipe the seal surfaces with acetone soaked wipes.

### Plug in breakout box connectors (sample must be grounded) to the fridge. Check sample wiring connections and measure the resistances of each wire as well as shorting to the chassis. Also use this time to double check the setup.

### Put sample device into the sample holder.

### Put radiation shield onto insert (optional):

##### Remove bottom cap, ensure the end piece is not touching sides of the rad. shield.

##### Cover any gaps of the radiation shield to the fridge with aluminium tape (not mylar).

### Remove Sorb and heat it up for a few minutes with a heat gun to remove any water condensation.

#### Be gentle – if copper mesh is changing colour or smoking, then it’s too hot. Replace Sorb when done (Sorb should be warm when IVC goes on).

### Ensure the IVC can is clean. Make indium seal.

#### Cut a piece of Indium wire 30cm long.

#### Clean/straighten with a methanol soaked kimwipe.

#### Vacuum grease Indium wire (very small amount).

#### Wrap In wire around can-side seal/flange. Use a scalpel to close the seal and trim off any free ends.

### Slide IVC onto the insert:

#### Align cutouts, take care not to damage pickup tube and snow trap.  Take care not to knock the rotator (if equipped).

#### Bring IVC into contact with the flange and finger-tighten all the screws.

#### Use an M4 size Allen key to tighten bolts one by one, rotating around the seal to close the seal evenly.

### Connect the external turbo pump to the IVC and pump the IVC out for 1-2 hours:

#### Use KF 16 elbow and KF 16-25 adaptor to connect the IVC port to the pumping line.

#### Check that the IVC valve is open.

#### Start the pump in rotary mode

###### use arrows < > to navigate to mode 023:Motor TMP off mode and pump down to around 5x10-1 mbar. This takes a few minutes.

#### Switch pump to turbo mode:

##### After around 1hr the pressure should be around 1.4x10-3mbar (good enough for some applications)

##### If turbopump is left on overnight, the vacuum will get down to the x10-5mbar range.

#### When at the right pressure, switch off and disconnect the turbo pump, to do this:

##### Shut IVC valve

##### Power off turbopump

##### Wait ~10mins, until the pumping line is vented

##### Disconnect the turbo pump

##### Remove KF-16 elbow and the adaptor off

### Introduce half a neck-full of He exchange gas from an orange bladder into the IVC. To do this:

### Pump  the 1KPot out:

#### Connect 1Kpot line to the 1Kpot S/V on the insert

##### (may need small extension line if 1Kpot line is not long enough).

#### Start He4 rotary pump (or turbo in roughing mode)

#### Open the blue 1KPot valve on the ICEcube, pressure on the 1KPot line P4 should drop.

#### Pump out the 1Kpot (1-2 min):

##### Check that the N/V is shut

##### Open the 1Kpot S/V on the insert

##### Shut the 1Kpot S/V (once vacuum is established P4 ~ x10-2mbar)

##### Shut the pump off when complete

### Flush 1KPot with Helium Gas:

#### Attach an orange bladder full of He to the 1Kpot S/V on the insert (flush fittings)

#### Open 1KPot S/V (let He gas into the 1KPot)

#### Open N/V slightly until you see the bubbles at the end of the pick-up tube (use beaker with methanol)

##### Note: Keep the bladder on with He gas flowing out of pickup-tube during the loading process, until insert is half-way down cryostat (prevents air from getting in).

### Introduce Exchange Gas

##### Check IVC valve is shut.

##### Fill orange bladder with Helium gas and clamp the neck of the bladder with a plastic clamp to keep the He gas trapped in the bladder.

##### Attach He orange bladder via a KF16 tree directly to the IVC port, flush the fitting & clamp in place.

##### Next move the plastic clamp on the neck of the bladder about halfway down and close the plastic clamp. This traps ½ a neckful of He between the bladder and the IVC valve.  Be sure the clamp is shut properly, so excess He is not introduced.

##### Open IVC valve to suck the ½ neckful of He gas into the IVC. Observe the lower half of the bladder’s neck flatten completely as it is now also under vacuum.

##### Shut IVC valve, disconnect the bladder from the IVC port. Add blanking plate and O-ring to the IVC to ensure it is completely sealed.

## Loading the Insert

### Prepare to load the fridge insert into the Cryostat - attach metal lifting posts to top of fridge insert and put the fridge on the hoist.

### Ensure seal around the O-rings and O-rings themselves are clean between:

##### Insert and cryostat

##### Insert and sliding seal

### Put sliding seal over the insert. Match the alignment markers and bolt it on.

### Attach rubber hose to exhaust of sliding seal (X-mas tree fitting).

### Record Helium meter value and level meter value

### Set level meter to fast mode

### Load insert with sliding seal into cryostat :

#### Set cold plate heater to max power on the white control box

##### This creates large overpressure in the M.B so no air gets in

##### Ensure the D-9pin connector to the “heater and sensor” port on the white control box is plugged in. Raise the temperature of the heater to the maximum possible - 1285mW (press escape, then press A+B until reach max. Heater mode, then press OK).

#### Take off L-plates on cryostat bung and remove cryostat bung

#### Carefully lift and slide insert into cryostat up to the black marker

##### (ie. sliding seal just covers cryostat, so there is a good seal where no air can get into the cryostat.

#### Wait for trapped air in sliding seal to escape through the rubber hose.

#### Attach rubber hose from sliding seal to the “To Sliding Seal” S/V on the M.B. exhaust line (once He4 gas is coming out of the rubber hose).

#### Open the “To Sliding Seal” S/V.

#### Open blue ball valve on the exhaust line to the recovery – meter should start ticking.

#### Close red striped ball valve to force the He gas through the sliding seal to reach recovery

#### Stop cold plate heater

##### Disconnect the cable from the cold head controller box, or just press escape.

#### Lower the insert slowly into the cryostat (so the second last digit on the recovery meter, ticks over every 5 seconds).

##### Note: during the last leg of the insertion process, need to use heat gun and paper towels to keep sliding seal dry when loading in.

##### IF insert doesn’t sit straight, it could be due to air pressure build-up in the M.B. – open red ball valve slightly to release gas pressure in M.B. , prop it up and closed the red ball valve when stable.

### When  insert is loaded half-way in – Pump and Flush the 1KPot:

#### Shut N/V

#### Shut 1kPot S/V

#### Disconnect the orange bladder of He from the 1KPot flange

#### Connect 1KPot pumping line (and small extension line if needed)

#### Start He4 rotary pump

#### Initial check - open blue 1KPot valve on ICEcube (pressure on 1KPot line P4 should drop to x10-2mbar). Then do pump & flush cycle.

##### To pump:

###### Shut the N/V (between MB and 1KPot, stops He4 from flowing into 1KPot)

###### Open the 1Kpot S/V to the insert (vacuum out 1KPot, P4 drops to x10-2mbar range)

##### To flush:

###### Shut the 1Kpot S/V to the insert (stop sucking out the He4)

###### Open the N/V (let He4 from M.B. into the 1Kpot, P4 should rise up to 1000mbar)

#### Repeat pump and flush cycle 2-3 times

#### Pump out 1KPot so it is at vacuum

#### Close 1KPot S/V on insert and blue 1KPot valve on ICEcube.

#### Disconnect 1KPot line (take away small extension line if added)

### When fridge fully in cryostat, bolt insert down to the cryostat with insulated plastic washer wrapped M4 screws.

### Close blue ball valve on the M.B. exhaust line and open red striped ball valve.

### Close “To Sliding Seal” S/V.

### IF unplug cable to white control box - reconnect the cable from the cold head controller box, and exit the heater override mode the white control box.

### Set level meter to slow mode.

### Connect the diagnostic cable (to lower fischer plug with yellow label) and note down the readings of gas meter and level meter. Turn AVS47-B on and monitor the Sorb (breakout box’s DB25 goes into pre-amplifier port at back of AVS).

### Once all ok, check the 1kPot one more time:

#### Connect 1KPot line.

#### Open blue 1KPot valve on ICEcube, to pump out the line b/w insert and ICEcube.

#### Open 1Kpot S/V on the insert to pump out the 1Kpot itself.

#### When 1Kpot pressure P4 drops to x10-2mbar, close the blue 1Kpot valve on ICEcube.

#### Open N/V to flush the 1Kpot

#### When P4 ~ 1000mbar, shut N/V.

#### Open blue 1Kpot valve to pump out the 1Kpot

#### When 1Kpot at vacuum (P4 ~10-2mbar)

#### Flush the pot  again and  leave fridge in this ending state (close 1KPot valves):

##### Close 1kpot N/V

##### Close 1Kpot S/V on insert

##### Close blue 1Kpot valve on ICEcube

### Get ready to pump out the fridge lines (OR N2 Cold Traps first if these are warm).

## Pump Out N2 Cold Traps

Note: This step can be done any time before the fridge lines (still and condenser) between the insert and ICEcube are pumped out.

### Turn on He4 rotary pump.

### Open light blue 1KPot valve

##### P4 on 1KPot line P4 drops to x10-2mbar range.

##### Be sure 1K S/V is shut, if connected.

### Open V14 on ICEcube

##### should be at vacuum, so almost no change in pressure

### Open V13 to N2 traps

##### should see a rise in pressure to 1mbar because of the traps, before it eventually drops down to 2x10-1mbar.

### As mixing lines are being pumped out, heat N2 traps with heat gun to boil off any trapped gas. Apply heat for ~5mins, until hot to touch.

### To stop pumping out N2 cold traps:

#### Close V13

#### Close V14

#### Close blue 1Kpot valve on ICEcube

#### Stop He4 pump

## Pump Out Condenser and Still Lines

### Connect He3 condenser and Still lines from ICEcube to insert.

### Connect turbopump to S/V labelled “He3” using a short tube and adaptors.

### Start turbo pump in roughing mode

##### (because there’s air/atm. pressure in both lines)

##### First use arrows < > to navigate to mode 023:Motor TMP off (press both arrows at the same time to confirm your selection)

### Open “He3” S/V to the turbo pump slowly and watch P2 on the Still line drop. (Ensure turbo pump pressure doesn’t rise above 200mbar during process).

### Open V9 (when P2 ~1.1mbar)

### Open V10 (slowly) to condenser.

#### First pressure on Turbo will rise to 22mbar then drop down

#### Watch pressure P1 drop from 1atm down to 0mbar

### Switch to turbo mode when P2 ~ 5x10-1mbar

### Wait for ~ 2hrs to complete.

### When finished, close V10 and V9 and the small “He3” S/V on cross piece to the turbo pump.

#### Now P2 ~ 3x10-3mbar and holding the vacuum, P1 = 0mbar.

### Stop turbopump (but do not disconnect – leave it for ~10mins to decelerate and vent back until it is around atmospheric pressure before disconnecting).

### When vented, disconnect turbopump and blank off He3 flange.

### If pumping lines overnight  turn the sorb heater on to bring the system to 4k. When ready to put the mixture in turn the sorb heater off.

# Transfer Liquid Helium (cool to 4.2K)

### Want to transfer LHe into the M.B. of the cryostat.

### If cooling the M.B. down from room temperature, log magnet resistance values along the way.

| Time | P M.B. (mBar) | PT2 Cmpr (() | Rmagnet (() | R1 (() | R2 (() | R3 (() |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 0:00 | 10 | 3611 | 39.2 | 241.5 | 269.5 | 262.1 | Start: He level ~0% |
| 1:30 | 7 | 23,837 | 0.3 | 1.241k | 1.199k | 1.077k | Finish: He level ~89.8% |

### Before starting: check if the small security valve on the 4He dewar is opened and check the pressure of the dewar. Don't forget to put back in the appropriate box the fittings and seals for the transfer.

### What you need:

#### Cryostat side siphon (with bung, seal nut, plastic washer – these are in the bottom drawer of the white sliding drawer cabinet labelled ‘fridge parts’)

#### Dewar side siphon with hose

#### Protection gloves

#### He level dip-stick

### Dip the Dewar (to check the amount of He Level)

#### Measure He level in the transport dewar with the dip stick. Close the vent valve. Remove bung piece of transport dewar and open transport valve. Measure He level by measuring distance between bottom of transport dewar and phase boundary (put a finger on membrane formed by rubber glove, He4 level found when vibration changes from low (liquid) to high frequency (gas)). Use meter ruler to measure height where vibration changes.

#### Repeat He level measurement again.

### Starting the LHe Transfer:

#### Note down the level meter and gas meter readings before the transfer

#### Set level meter on fast.

#### Place transport dewar close to cryostat and connect rubber hose from dewar vent (with red tape) to the ”To Transport Dewar” S/V on M.B. Exhaust.

#### Turn cold plate heater to max power on white control box

#### Check if the adaptor piece is attached to the transfer port of the cryostat

#### Open the transfer port on the cryostat and insert the cryostat side siphon:

##### Take off top bung of transfer port adaptor and insert cryostat siphon.

##### Put brass screw nut and the plastic o-ring on top-end of cryostat siphon, ready for coupling of transfer siphons.

##### Close manual recovery valve (red striped ball valve) and open the main blue valve. (so helium gas is forced out the siphon).

##### Insert cryostat siphon slowly all the way down into cryostat.  Once completely inside and white torch is coming out of siphon, close end with rubber bung.

##### Open manual recovery valve (red ball valve)

#### Check 4He4 dewar pressure.

#### If dewar almost empty (~20 cm), don't forget to attach extension tube to transport side siphon.

#### Insert transport side siphon into transport dewar:

##### Take off top bung of transport dewar and open top transport valve

##### Close vent valve and release valve on transport dewar

##### Insert siphon slowly (make sure that there is not much pressure build up in the dewar.

###### *If the pressure is too high, liquid might start flowing out of siphon before it is fully in, making it impossible to couple the siphons.

##### Pressure of He gas flowing out will increase and it will become colder and whiter. When white torch is coming out and siphons are possible to couple, take off bung from mating siphon and couple two siphons together quickly.

##### Screw seal nut tightly.

### Stop cold plate heater

##### (Disconnect the cable from the white controller box)

### Monitor pressure in transport dewar, He gas meter and He level meter. If not enough pressure in transport dewar, and transfer is too slow, you can:

#### Use a small rubber balloon (with bag) and connect it to the vent port on the transport dewar. Open vent valve and see balloon expanding, pump it by hand several times and close vent valve.

OR

#### Plug hose of He4 gas cylinder into He4 dewar vent valve (with red tape) – if necessary disconnect and bung rubber hose going to M.B. exhaust before attaching this. Open He4 gas cylinder valve to let gas flow into the dewar to increase the pressure, hence the transfer rate. Check pressure on the transport dewar, keep it around 2.5 psi (don't put more than 3 psi). Check that recovery meter rate is about 1 second place digit every 5 seconds. Check that the level meter changers by <1 % per reading.

### Stopping the transfer:

#### Release the pressure in the transport dewar:

##### Put a rubber hose onto transport vent valve (take off orange bladder in navy fabric if pumping dewar previously to transfer)

##### Open vent valve slightly to let He bleed through rubber hose and remove air (test pressure of He gas coming out by blocking other end of hose with thumb)

##### When air is displaced, attach other end of hose to the “To Transport Dewar” S/V flange on the M.B. exhaust line

##### Open vent valve on transport dewar fully

##### Open “To Transport Dewar” S/V slowly (ensure no large rush of pressure from gas)

#### Wear protection gloves, unscrew the screw nut and decouple two siphons.

#### Bung cryostat side siphon.

#### Pull out transport dewar siphon from transport dewar (try to keep it straight so it doesn’t bend)

##### Close vent valve and transport valve

##### Open release valve

##### (keep the small He4 dewar safety valve open)

OR

#### If using the liquefier - bung the liquefier transfer port.

#### Close main ball blue valve on exhaust.

#### Pull out cryostat siphon from cryostat

##### If the adaptor pulls out of transfer port place it back with bung fitted.

#### Reconnect cable from the cold head controller box, and exit heater override mode in the white control box

#### Set He level to slow mode.

#### Record gas meter and level meter again after transfer.

#### Dip the LHe transport dewar again

# Cool Fridge to Base Temperature from being Parked at 4K

3 main stages to the process:

- Cool 1KPot

- Condense He3/He4 mixture

- Circulate He3/He4 mixture

Procedure:

### IF fridge is parked at 4K for some time, may have heat leakage – shown by a slight rise in 1Kpot temperature from 4.2K. So first heat the Sorb to Tsorb = 30K, for around 30-45mins (takes around 30mA of current).

#### Wait for the Sorb temperature to go down to ~4K, RMC=2.78 kOhm

#### Turn the Sorb Heater off.

### Fill the N2 traps.

### Apply max. power to the cold plate heater via the white control box on the ICEcube.

### Insert He Trap into cryostat (done anytime before condensation, when not pumping the lines)

#### Stop pumping condenser & still lines – shut V10 and V9 (all valves should be shut)

#### Shut He3 valve to turbo pump

#### Set Cold Plate heater to maximum to create an overpressure in the cryostat

#### Slowly insert He4 Trap, so not too much LHe is boiled off

#### When halfway through inserting He4 trap – stop heating cold plate

#### When done ensure seal around He4 trap to cryostat is tight

### Cool 1KPot

#### At 4.2K, 1KPot temperature (2.54k( – AVS Ch1)

#### Turn He4 rotary pump ON

#### Preliminary check that 1Kpot is working:

##### Check T1Kpot = 2.52 kΩ.

##### Open blue 1KPot valve on ICEcube,  then open the 1KPot S/V on the insert. to pump out the 1KPot (P4 should drop <1mbar (ie. x10-2mbar) when the 1KPot is empty.

##### Shut blue 1KPot valve.

##### Preliminary flush of 1KPot:

###### Open N/V between pickup-tube and 1Kpot fully for 3 seconds (allow LHe4 to enter into the 1Kpot)

###### Watch 1Kpot temperature drop, wait for it to reach ~1.6K

#### Then pump out 1Kpot:

##### Open 1KPot blue valve on ICEcube to pump on 1Kpot line, wait for P4 on 1Kpot line to drop to x10-2mbar.

##### Open 1Kpot S/V on fridge insert to pump out 1Kpot, 1Kpot line pressure P4 should drop <1mbar (ie. x10-2mbar) when the Pot is empty.

#### To cool 1Kpot:

##### Open N/V b/w 1Kpot and M.B. to feed LHe4 into 1Kpot, then close it again (open from marker 90o anticlockwise to 40o and back to 0o).

##### Set N/V so that P4 stays ~12mbar

##### Wait for 1Kpot temperature to drop to ~1.6K (3.42k()

#### Ready to start condensation when T1kpot = 1.626K, P4 = 2.5mbar,  P2 = 5x10-4mbar

### Condense the He3/He4 Mixture

#### Open both green He3 dump in & dump out valves ontop of the dump (a grey cylinder)

#### Switch on He3 rotary pump to circulate mixture through ICEcube.

#### Watch P3 pressure gauge at back of He3 pump rise. Also P2 (pressure on Still line attached to front of He3 pump) will rise a bit higher.

#### Open lines to fridge insert – open Still gate valve and Condenser S/V

#### Open V9, to let mixture fill line in between the condensing line and still line.

#### Open V10 (slowly) to condensing line

#### Open V6 to the dumps (closest to back of rotary pump) – mixture flows into the lines

#### Record starting P3 value in back of log book.

#### Open V5 (to N2 cold trap Input)

#### Open V4 (slowly) to N2 cold trap output

##### Monitor the following values:

###### P1 - should remain <200mbar during the operation (control this pressure via V4)

###### P3 ~500mbar at the start, should be dropping

###### P2 = P4 = 1000mbar

###### T1KPot – should remain <2.2K (ideally ~1.6K) ie. >3.1kΩ.

##### P1 (condenser) should rise slowly and P3 (back of He3 pump) drop until both pressures are equal and equilibrium between the two parts is reached – this takes >30 mins.

###### **IF in the event that P1 > 200mbar, check 1Kpot temperature – if T1KPot > 2.2K, it means that 1KPot is empty of LHe4 and fridge is warming up rather than cooling!!

Eg. P1 = 220mbar, P3 = 250mbar , P2 = 1000mbar , P4 = 3mbar

If this happens, shut V4 (stop mixture flowing from N2 traps into rest of fridge lines and raising pressure and temperature)

Open N/V a bit more (may need to tweak this a bit). As T1Kpot drops, close N/V - T1kpot should drop back to ~1.8K.

Once back to normal when readings meet the previous conditions of operation (say ~ : P1 = 200mbar, P3 = 250mbar, P2=P4=1000mbar, T1kpot = 1.8K)

##### When P1~P3 ~200mbar & P2=P4=1000mbar, open V4 fully without P1 rising as pressures in both lines have equalized

#### Temperature of 1Kpot now drops slowly (*note we want this slightly higher than the standard 1.5K, say ~1.6K)

#### Wait until both P1 and P3 drop below 100mbar (takes ~20mins)

### Circulate the Mixture

#### Close V9 to seal direct link between condenser and still lines

#### Close V6 to dumps

#### Open V8 to pump out everything else from dumps

#### Open NW40 (slowly) to He3 rotary pump – note that NW40 is very sensitive at the start, though towards the end can open faster/more.

#### Watch P3 at back of pump rise slowly & since P3 is rising, so will P1.

##### Note: P3 on back of He3 pump, should be kept below 180 – 200 mbar during the whole process...

##### Also monitor TMC~1.8K. We need this to drop to ~100mK for phase separation of the mixture to begin.

#### Monitor P2 – when this drops to ~50mbar and P3~130, not much difference between the two sides of the lines so can open up NW40 fully.

#### At end of circulation, T1kpot~1.73K, P1<P3<100mbar, P2 = 6.2x10-1mbar, P4 = 17mbar

#### Close V8. For extra safety it is possible to open V6 but it is not necessary because there is an overpressure valve in parallel with V6.

#### Turn on Still heater - plug it to diagnostic box pins 21 and 22 (set current to 3.6mA for 6mW of power output)

###### This helps to promote flow from the dilution phase liquid into He3 gas back up through the still pumping line into the dumps.

#### At base fridge is around: TMC~37mK, T1kpot~1.6K, P1~50mbar, P3~70mbar, P2~0.35mbar, P4~9mbar.

# Lift Fridge out with Cryostat Cold

There are 3 main stages to lifting out the fridge:

- Stop circulation of mixture

- Stop 1KPot (condensation of mixture)

- Heat the fridge up to pull the mixture out, using the He3 rotary pump

### Before starting: You will need protection gloves, a ladder, allen keys, a heat gun, plastic caps for pumping lines.

### Disconnect sample from instruments and ground it.

### Check dump taps open.

### Close V5, to stop circulation.

### Open V9 to equalize pressures in the two lines (condenser and still).

### Open V6 to provide a path to dumps (if not already so).

### Close V8 so mixture is forced into dumps via the back of the He3 pump through V6 (the mixture will be pulled out using the He3 rotary pump)

### Shut 1KPot by closing N/V

### Note that at this stage, most of the mixture is still in liquid form

### Next start heating up the fridge:

#### Turn Sorb heater on [connect Sorb heater to breakout box pins 19, 20 & set i = 30mA].

#### Pump out 1KPot till it is empty

#### Turn M.C. heater on [TS-530A BNC out to pin 23, gnd cap to pin 24]. Use TS-530A controller values Bias = 20%, Range = 100mW, readout set to power [pin 23 to TS-530A, pin 24 – grounded]

#### Turn Still heater on [connect Still heater to pins breakout box 21, 22 & set I = 7mA]

### Finishing values (after waiting for a while):

##### P2 = 7x10-4mbar, P4 = 3.4x10-2mbar, P3 = 600mbar, P1 = 0mbar

##### TMC = 2.38kΩ, T1kpot = 2.34kΩ, Tsorb = 0.64 kΩ, Tstill = 2.5kΩ, TCP = 2.60 kΩ

### To pull mixture out, close and open the NW40 S/V:

#### First Close NW40 valve

#### Start to open NW40 slowly, such that P3 rises

#### Leave this open for 1hr, whilst waiting for P3 to reach 600mbar

#### When done pressures are: P2 = 8.0x10-4mbar, P4 = 2.1x10-2mbar, P1 = 0, P3 = 600mbar

### Record final P3 value in back of log book.

### Switch off all heaters (Sorb, Still & MC).

### Close V4, V9, V10, NW40 and V6.

### Close dump valves.

### Shut Still, Condenser and 1KPot valves on fridge insert.

### Shut blue 1KPot valve on the ICEcube.

### Record P3, He level, recovery gas meter, room temperature and pump temperatures in the log book. All valves at this stage should be closed. He level (set to fast mode) and recovery gas meter.

### Turn pumps off (He3 first then He4)

### Turn all equipment (diagnostics) back on.

### Set cold plate heater power to max. 1285mW on white control box

### Take He trap out of cryostat carefully – untighten nut to lift out, make sure to keep it straight as it is being lifted out, and put cryostat bung back on immediately when trap is out. Then warm the He trap up with a heat gun to room temperature.

### Disconnect fridge lines to insert in order (1Kpot, Still, Condenser) – ensure their valves are closed first.

#### Blank off fridge line ends to ICEcube using the same plastic clamps and O-rings used b/w the lines and the cryostat. Put plastic caps over flanges on the insert.

#### Now P1 = P2 = P4 = 1000mbar (coz at ATM) & P3 ≈ 500mbar

### Prepare to lift out insert:

#### Record He gas meter and He level.

#### Disconnect diagnostics cable.

#### Attach metal lifting posts to top of fridge insert. Tie sample boxes with thick wire and attach hoist.

#### Undo plastic insulated screws holding sliding seal to cryostat.

#### Switch the white controller box to monitor pressure P in the cryostat

#### Check fridge is clear of all cables and pipes, so nothing will catch when lifting.

#### Start lifting the fridge  - monitor pressure in the cryostat, want positive P in the M.B. so no air gets sucked in during lifting

### When fridge insert is almost out ie. Reached the black mark on the sliding seal:

#### Check “To Sliding Seal” S/V on M.B. exhaust line is closed.

#### Get ready to lift remainder of insert out faster (need 2 people – one to lift, one to guide). First wait for a bit for pressure to build in M.B., then lift the insert all the way up carefully and slide it away from the cryostat.

### Put back top bung in the cryostat and tighten it with special L shaped corners and short M4 screws.

### Disconnect rubber bypass line from sliding seal.

### Record values after liftout – He level (set back to slow), He recovery Gas meter, PT2, Pressure in MB and Heater.

### Remove sliding seal carefully from insert - undo screws which hold sliding seal and tilting insert slightly as necessary.

### Stop heating cold plate (overpressure is no longer needed) via. white control box (wait for a while before performing this step).

### Mount fridge insert on mobile fridge support stand. Use a G-clamp and strip of perspex to lock it in position securely. Be careful the stand doesn't roll away from you.

### Leave insert on support stand overnight for IVC and insert’s internals to warm up to room temperature before removing IVC.

### As fridge is left overnight, hook up a turbo-pump to fridge lines of the insert via a T-piece (Note – can go straight to turbo mode as the lines are still under vacuum and are cold) [*see step2 on page 1]

## Take off Inner Vacuum Chamber can and Access Sample

### Once at room temperature, there’s still a vacuum in the IVC.

### So first vent the IVC by opening the IVC valve (if there’s a blanking plate, remove it)

### Undo the screws from the IVC. Use two special jacking screws to jack off the IVC and break the indium seal.

### Slide down the IVC can carefully.

### Undo the screws for the radiation shield and remove that can.

### Now the sample is accessible.

# Cooling Cryostat from Room Temperature

## Cryostat Preparation

If cryostat hasn’t been used for a while and is sitting at room temperature (ie. Is filled with air), need to pump and flush:

- The M.B.

- The Lambda Plate (attach orange bladder)

- The 1KPot

Procedure:

### Prepare to Pump and Flush the M.B. of the Cryostat

#### Connect large exhaust line (cross piece) to M.B. exhaust port

#### Connect external turbo pump to M.B. valve on cryostat

#### Pump down M.B. until pressure is 5.1x10-1mbar

### Pump out the 1Kpot

#### Connect fridge lines to the 1KPot

#### Switch on He4 pump near 1KPot.

##### Watch the pressure gauge on the M.B. drop to 1.8x10-1mbar.

##### P4 (1Kpot gauge) should drop from 1000mbar to 1.8mbar

##### When open N/V (b/w 1Kpot and M.B) P4 drops further down to 1.6x10-1mbar.

### Pump and Flush with N2 gas (x3)

#### Pump out the M.B. (via turbo pump)

##### Note: Since there’s a lot of volume in the M.B., the valve to pump the M.B. should be left mostly closed, until M.B. pressure gauge on the exhaust line starts to drop. Only then open up the valve to the M.B. fully to pump out the whole thing (otherwise damage pump).

##### Ensure pump out everything (M.B., Lambda Plate, 1Kpot) then close valve connections between Lambda Plate and 1KPot to the M.B.

#### To flush M.B. with N2 gas:

##### Ensure blue ball valve is shut, attach N2 gas dewar to the end of the M.B. exhaust line.

##### Open blue ball valve to N2 gas dewar to flush M.B. and monitor M.B. pressure on white control box (on side of ICEcube)

###### Note: With the current setup, the white oxford box readout shows 64mbar for M.B. pressure, this is only an overpressure reading.

##### When M.B. pressure stops rising ~64mbar max., close blue ball valve to stop flushing with N2.

#### To pump:

##### Slowly open valve from turbo pump to M.B. pump, to pump out N2 gas. Wait till M.B. pressure drops to ~1mbar

###### Note: whilst we do this, the 1Kpot is being pumped through the pickup tube between 1Kpot and the M.B.

#### Once this pressure is reached,

##### open valve to the lambda plate

##### Open N/V

##### Turn on He4 pump to pump out 1KPot completely

##### Open 1KPot S/V on the insert to the 1KPot

###### Eventually M.B. pressure should drop to 3.6x10-1mbar and 1Kpot pressure P4 should drop to 5.8x10-2 mbar.

#### Repeat this pump and flush cycle a total of 3 times

#### Ensure the M.B. is completely pumped out before begin flushing with He

### Pump and Flush with He gas (x2)

#### Flush with He Gas

##### Check blue ball valve on M.B. exhaust line is closed, remove N2 inlet from end of M.B. exhaust line and replace with He inlet (flush fitting first)

##### Close S/V to M.B. turbo pump (stop pumping)

##### Open blue ball valve to let He into M.B.

###### M.B. pressure gauge on exhaust line >100mbar, and P4 on 1Kpot line slowly rises to 1.1x10-1mbar and higher.

##### Open valve to orange balloon attached to lambda plate (Balloon fills with He and expands)

##### Then flush N/V (by opening it, letting He through, then closing it)

##### Open 1KPot S/V on the insert to allow He to flow through the 1Kpot and be pumped out via the He4 pump, then close the 1Kpot S/V on the insert.

#### Pump out He Gas:

##### Close blue ball valve to stop He4 gas flowing into M.B.

##### Check all valves are closed (lambda plate, N/V, 1Kpot valve)

##### Open M.B. S/V (slowly) to pump out He4 in M.B.

###### M.B. pressure gauge on exhaust line should drop <100mbar

###### *can also hear change in pitch of M.B. pump

##### Open valve to lamba plate balloon (which deflates as He is pumped out), then close the valve.

##### Open N/V to pump out, then close it.

##### Open 1Kpot valve to pump out 1Kpot. Watch P4 drop to 5.4x10-2mbar. Then close 1Kpot valve.

##### Close S/V to M.B. pump.

#### Repeat steps a & b one more time for the second He pump & flush cycle.

### Once He Pump & Flush cycle is completed, check all values are closed

#### Turn off turbo pump attached to M.B. exhaust line

#### Turn off He4 rotary pump

## Compressor Stage

### Turn Compressor on (room next door).

### Takes a total of ~3days to drop down to the right temperature

### To monitor the temperature as fridge is cooling, read out magnet resistances as well as the resistance of the cold head PT2.

| Time | P M.B. (mBar) | PT2 Cmpr (() | Rmagnet (() | R1 (() | R2 (() | R3 (() |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | 257 - 279 | 47.4 | 214.2 | 235.9 | 241.7 | Room Temp. |
|  | 2 (fixed) | 3611 | 40.1 | 234.1 | 260.6 | 260.9 | End of compressor run. |

### At the end of the compressor run, the readings from the white compressor control box should be around:

P(M.B.) = 17 mbar	Heater = 347 mW	PT2 = 5499 Ω

### From diagnostics, temperatures of the fridge parts are:

Sorb 	(Ch0)   =  1.38 kΩ

1KPot 	(Ch1)   =  2.41 kΩ

Still 	(Ch2)   =  2.44 kΩ

Cold Plate   (Ch3)  =  2.50 kΩ

Mixing Chamber  (Ch4)  =  2.53 kΩ  (Femto Power)

### Once set, prepare to do LHe transfer.

# Maintenance during Operation at Base Temperature

### Connect transport dewar to recondense vapour (if find there’s too much He boil off)

#### Connect orange bladder to transport dewar to create positive over pressure.

#### Connect rubber hose from transport dewar nozzle to M.B. S/V labelled “To Transport Dewar”, using an X-mas tree fitting.

#### Flush the fittings.

#### Open vent valve on transport dewar.

#### Open the “To Transport Dewar” S/V.

### Disconnecting the Transport Dewar (if He  level in dewar is getting too low)

#### Shut S/V labelled “To Transport Dewar” on M.B..

#### Shut vent valve on transport dewar.

#### Disconnect rubber hose from the M.B. exhaust line.

### Clean the Cold Head

(Theoretically the K100 system should have zero boil off, but sometimes we lose 2% per day on the level meter- this could potentially be caused by a dirty cold head)

Note: Boil off occurs when the level meter drops but the He recovery meter doesn’t move

#### He losses in the fridge that’s not being recovered via the recovery line.

#### Plug in the D-9pin cable to the heater & sensor port on the white control box attached to the IceCube.

#### Set cold plate heater power to maximum 1285mW.

#### Turn compressor off.

#### Let it be for ~2hrs.

#### Turn off cold plate heater by unplugging the D-9pin cable and turn the compressor back on.

# Nomenclature

M.B. = Main Bath

S/V = speedivalve

N/V = needlevalve