Operation procedure Edwards Auto 306 Evaporator with Turbo

## Loading

- Normally evaporator should be left pumping (three way switch is in ON position and the Turbo saver switch is in 60% position).

- Check the pressure in the evaporator.

- Check that the sample holder is outside – no samples are inside evaporator.

- Turn the three way switch into VENT position and wait until the pressure goes up to >1.2x103 mbar.

- Check if you can open the chamber door by trying to release the door lock.

- If the door can be open, turn  the three way switch into OFF position.

- Open the yellow extraction valve at the back of the chamber.

- Release the lock and open the door.

- Put the plastic hood and remove three bottom metal shield plates.

- Close the gate valve by switching the gate valve switch which is located inside the cabinet underneath the vacuum chamber. Observe the gate valve closing.

- Put the metals in the boats according to convention. The boats are mounted on the rotation stage, which can be moved by turning the manual knob below the door.

- Turn the knob clockwise so the required boat is touching the graphite current lead.

- Put the sample holder into the BNC type slot on the ROTATILT system.

- Check that the manual shutter SS1 is closed.

- Set the appropriate angle by manually tilting the sample holder (test rotation if required).

- Put the metal shield plates back in place except the bottom one.

- Open the gate valve by switching the gate valve switch located inside the cabinet underneath the vacuum chamber. Observe the gate valve opening.

- Put the bottom metal shield plate.

- Remove the plastic hood at put it on the wall.

- Shut and lock the door.

- Close the yellow extraction valve at the back of the vacuum chamber.

- Turn the three way switch into ON position. Watch the controller screen 340 displays pressure going down.

- If fast pumping is required set Turbo saver switch to OFF position. Otherwise leave it at the 60% position.

# Deposition and Unloading

- Record chamber pressure in the log book.

- Switch the Turbo saver switch  into 60% position

- Turn on the Inficon Xtal monitor at the back of the unit.

- Check the Xtal value and write it down in the log book (if crystal life <20%, replacement needed).

- Depending on the evaporation material, select required FILM number.

- Set the rotation speed if required.

- Check that the manual shutter SS1 is closed

- Check that the power knob is in zero position.

- Turn the power switch to LT position (see the power LED goes green)

- Start slowly ramping up current by turning the manual power knob.

- Monitor the rate on the crystal monitor controller and the pressure in the chamber. When ready to evaporate, zero the Xtal monitor and open SS1 shutter simultaneously.

- Control the rate by turning the power knob.

- When the required thickness is achieved, close SS1 shutter.

- Ramp the current back to zero by turning the power knob.

- Switch the power switch from LT to OFF (middle) position.

- Turn the sample holder rotation back to zero.

- Wait for 2 mins to cool.

- If the second metal should be deposited turn the boat rotation stage (manual knob below the door) so that the appropriate boat touches the graphite current lead and repeat steps 5 to 17.

- Write down in the logbook rates and thicknesses of metals you have deposited.

- Turn off the Inficon Xtal monitor at the back of the unit.

- To take the sample out, turn  the three way switch into VENT position and wait until the pressure goes up to >1.2x103 mbar.

- Repeat Loading section steps 5 to 10.

- Remove the sample holder from the vacuum chamber.

- Repeat Loading section steps 16-23.

In case of any problem with the evaporator or if you do not know what to do contact Oleh on x55203