### STANDARD OPERATING PROCEDURE

ULVAC MILA-3000

Annealer

Ref: QED

Date:22/03/06

Review date: 17/04/08

Emergency Procedures and Shut down: Close all gas line valves including main gas bottles. Turn off power to annealer. Close water valves. Notify Custodian.

Description: An infrared lamp is used to rapidly heat a carbon susceptor and a sample contained inside a quartz chamber up to a maximum of 1200C. Air, argon, or a forming gas consisting of a maximum of 5% hydrogen in nitrogen may be used during the process. Volatile waste products are removed to the laboratory exhaust system.

Authorised Users: Oleh Klochan

Authorisation provided by: Oleh Klochan

Custodian of equipment: Alex Hamilton

Training required for task provided by: Oleh Klochan

Supervision required for task: Not required if trained.

Responsibility for review of this SOP: Jack Cochrane.

Resources:

Equipment: Annealer and connected nitrogen, argon and 15% hydrogen/nitrogen gas.

Substances: Samples, nitrogen, argon and 15% hydrogen/nitrogen gas.

PPE: Gloves/eye goggles

Ref Risk Assessment: See Risk Assessment Folder in Lab

Clean up and Waste Disposal: Waste solvents to be placed in Solvent Waste Disposal Drum. Waste glass to be placed in “Glass Waste” drum. Gloves to be placed in “Used gloves” bin

Control measures: Refer to Risk Assessment____

Key Hazards: Compressed argon gas. Compressed and flammable 15% hydrogen/nitrogen gas. Hot surfaces. Flammable gas.

Risks from Hazard: Refer to Risk Assessment.

Control Measure including PPE: Refer to Risk Assessment.

Note: Step by step instructions over page.

Rapid Thermal Annealer – Instructions

## Annealing/cooling gas setup

The current gas setup is based on the following operation steps:

Step 1. Venting the chamber before allowing annealing gas. For venting purposes, nitrogen gas line is used. Nitrogen is flowing through the right hand side 30 l/min flow-meter (by default rate is set on max ~17l/min). This process is imperative when forming gas is used for annealing in order to displace the oxygen from the chamber.

Step2. Venting/annealing the chamber with the annealing gas. Pressurized Ar or forming gas used as the annealing gas. It flows through the left hand side 1 l/min flowmeter. In case of forming gas, it should be diluted to =<5 % (non flammable) from 15 % (highly flammable) mixture in the gas cylinder. For that purpose the middle 1 l/min flow meter is used to allow nitrogen from nitrogen line to mix with the forming gas. Adjust the left hand side regulator to 0.2 l/min (undiluted forming gas) and the middle regulator to 0.4 l/min (nitrogen gas) giving total flow of 0.6 l/min of diluted (~5%) forming gas. In case of Ar gas adjust left hand side flow meter to required flow rate <0.6 l/min and keep the middle flow meter (nitrogen gas) shut.

Step3. Cooling the chamber/sample after anneal. For cooling nitrogen gas is used. Nitrogen is flowing through the right hand side 30 l/min flow-meter (by default rate is set on max ~25l/min).

NB Use the combustible gas leak detector each time the forming gas bottle has been replaced or gas setup is modified. SHOULD be done only by maintenance personnel.

## Start up

- Choose the annealing gas by setting the annealing gas valve (behind the laptop) into Ar (argon) or F (forming gas) position

- Open the main gas tap on the forming gas or argon cylinder

- Check the low- pressure regulator on the cylinder is set to 100 kPa

- Check that the three-way valve is in OFF position (above the flowmeters)

- Open (fully) the main nitrogen valve near the fume cabinet

- Open the annealing gas valve at the back of the annealer

- Check if the cooling water inlet tap and outlet tap are opened (behind the Edwards evaporator)

- Check if the 20A power point is turned ON (C.B. 1.15)

- Turn on the power switch on the right-hand side panel of the annealer

## Loading/Unloading

- Undo two bolts holding the sample moving flange and carefully slide it out

- Don’t touch the quartz sample holder with bare hands

- Put/take  your samples on/from  the carbon susceptor

- Slide the sample moving flange in place and tighten two bolts manually tight.

## Operation

- Vent the chamber with nitrogen gas (3 min at max flow) by turning the three-way valve into “N2 flush” position (flow rate ~17 l/min on right hand side 30 l/min nitrogen flow meter). Refer to Annealing/cooling gas setup section.

- Vent with annealing (forming or argon) gas for 2 min by turning the three-way valve into “gas” position. For forming gas (flow rate 0.2 l/min on left hand side flow-meter  + 0.4 l/min middle flow meter). For argon (flow rate <0.6 l/min on left hand side flow-meter + middle flow meter is shut). Refer to Annealing/cooling gas setup section.

- Press the HEAT button ON on the front panel of annealer (the LED should light on indicating that the power can be applied to the IR lamp)

- Choose the annealing pattern (preset) by pressing PTN button on the temperature controller front panel.

- Press and hold for about 2-3s RUN button on the temperature controller front panel (program starts)

- When the anneal has finished, to cool the sample turn the three way valve immediately into nitrogen “N2 flush” position. Refer to Annealing/cooling gas setup section.

- Shut the left hand side and right hand side flow-meters.

- Allow the system to cool-down for 10 min before removing the sample from the chamber. Sample temperature should not exceed 35 (C.

- Turn the three-way valve into OFF position

- Press the HEAT  button off on the front panel of the annealer (the LED should go off)

- To unload the sample follow the Loading/Unloading Section

#### Shutdown

- Check that the HEAT button on the front panel of the annealer is off  (the LED should go off)

- Check that the three-way valve is in OFF position

- Close the valve on the annealing gas line at the back of the annealer

- Shut the main gas tap on the corresponding cylinder

- Shut the main nitrogen valve near the fume cabinet

- Turn off the power by pressing the power button off on the right side of the annealer

- Switch off the water flow by closing inlet and outlet valves on the cooling water line (behind the Edwards evaporator)

- Clean up the area

- Notify responsible people if any problems occurred and fill the logbook

##### Programming Omron Temperature controller

###### Refer to MILA-3000 manual in order to program the temperature controller.

###### Hazards

##### Burns

After the anneal the quartzware may still be several hundred degrees Centigrade when pulled immediately. Cool down at least for 10 min before pulling the sample moving flange out.

##### Explosion

Too-high inlet gas pressures may cause backpressure, possibly resulting in an explosion. Keep the pressure on the bottle regulator set to 100 kPa.

##### Electrocution

This danger exists with any electrical system. Repair or troubleshooting may be done only by maintenance personnel unless specified.

##### Fire

Thermal runaway may result for a number of reasons, perhaps causing a fire. The machine must be closely watched during operation.

15% mixture of hydrogen in nitrogen is highly flammable. Avoid heat, sparks, contact with oxygen.

##### Asphyxiation

May result if changes in the operating procedures are performed without approval from the custodian of equipment. Also known as strangulation.